## Slide 1: Alinity i STAT High Sensitivity Troponin-I (hsTnI_STAT)
May 24, 2023

## Slide 2: AGENDA
*Intended use and Safety information
Heart disease and myocardial infarction (MI)
Diagnosis of MI: importance of biomarkers
What is troponin and how is it used in suspected MI?
How are hs troponin assays defined and reported?
Limitations of contemporary troponin assays
Comparison of Abbott contemporary and high sensitivity troponin-I (hsTnI) assays
Interfering substances
Alinity i hsTnI_STAT performance and specifications
Alinity i STAT High Sensitivity Troponin-I
August 14, 2023
ADD-140731-USA-EN |

## Slide 3: Alinity i STAT High Sensitivity Troponin-I
  INTENDED USE
  The Alinity i STAT High Sensitivity Troponin-I assay is a chemiluminescent microparticle immunoassay (CMIA) used for the quantitative determination of cardiac troponin I (cTnI) in human plasma (lithium heparin) on the Alinity i System.
  The Alintiy i STAT High Sensitivity Troponin-I assay is to be used as an aid in the diagnosis of myocardial infarction (MI).
  IMPORTANT SAFETY INFORMATION
  Instructions must be carefully followed. Reliability of assay results cannot be guaranteed if there are any deviations from these instructions.
  For laboratory professional use only.
  For In Vitro Diagnostic Use
  Rx Only (For use by or on the order of a physician only)
  This product contains human-sourced and/or potentially infectious components. It is recommended that these reagents and human specimens be handled in accordance with the OSHA Standard on Bloodborne Pathogens.
Alinity i STAT High Sensitivity Troponin-I Package Insert: H14937R02
August 14, 2023
ADD-140731-USA-EN |

## Slide 4: Heart disease and MI
Coronary Heart Disease is the leading cause of death in the US1
Manifests most often in the form of Acute Coronary Syndrome (ACS): term used to describe Myocardial Infarctions (MI, heart attack) and unstable angina
Approximately 6-8 million people/yr present to Emergency Departments with suspected MI1
Only 8-15% are having an MI so in theory the majority could be discharged2
Heart Disease and Stroke Statistics-2019 Update: A Report From the American Heart Association. Benjamin EJ, et al. Circulation. 2019 Mar 5;139(10):e56-e528.
Sandoval et al, AJC 2017
August 14, 2023
ADD-140731-USA-EN |
4

## Slide 5: Diagnosis of MI: importance of biomarkers
However, fears of “missing” an MI mean patients are frequently kept for observation and testing costing billions of dollars annually1
DIAGNOSING MI IS NOT ALWAYS SIMPLE
Symptoms may not be typical even with MI
EKG may not show changes even with MI
Circulating biomarkers are thus critical in making the diagnosis of MI (“ruling in” MI) or excluding the diagnosis of MI (“ruling out” MI)2,3
Troponins are the only biomarkers recommended for this purpose2,3,4
About 25% of patients with MI are diagnosed from their EKG due to ST elevation or new left bundle branch block = STEMI, ST elevation MI
The other 75% with MI but not STEMI = “NSTEMI”, non ST elevation MI
People who do not have STEMI (= suspected NSTEMI) are discussed from this point on (the STEMI diagnosis is usually made without the need for biomarker results)
However that is only 25% of the 8-15% with suspected MI.
Storrow, AIM 2000
Amsterdam E et al, JACC 2014 (ACC AHA NSTEMI Guidelines)
Thygesen et al, JACC 2018 Fourth Universal Definition of Myocardial Infarction
Yau AA, Nguyendo LT, Lockett LL, Michaud E. Crit Pathw Cardiol. 2017 Dec;16(4):126-128
August 14, 2023
ADD-140731-USA-EN |

## Slide 6: What is troponin? How is it used in suspected MI?1
Troponin is a protein contained in the cardiac myocytes that facilitates myocyte contraction with three subunits: troponins C, T, and I: Troponin T and I are highly specific to cardiac muscle, I more so2
Levels become elevated (>99th percentile) when myocytes are deprived of oxygen = myocardial injury including (but not limited to) MI1
Today, detection of an elevated level of troponin (> the 99th percentile) with a rising and/or falling pattern in the appropriate clinical setting forms the cornerstone of the diagnosis of MI1
Notably, most apparently healthy individuals (not having an MI) have small but detectable concentrations of troponin, but at these low concentrations, only high-sensitivity (hs) assays can detect it
Thygesen et al, JACC 2018 Fourth Universal Definition of Myocardial Infarction
Wens et al Circ Genetics 2016
August 14, 2023
ADD-140731-USA-EN |

## Slide 7: Guidelines for troponin in suspected MI: AHA/ACC1
“high-sensitivity cTn is preferred because it allows rapid detection of myocardial injury and has increased diagnostic accuracy”
Gulati et al, 2021 AHA/ACC Clinical Practice Guideline
August 14, 2023
ADD-140731-USA-EN |
7

## Slide 8: Guidelines for troponin in suspected MI: FUDMI1
Thygesen et al, JACC 2018 Fourth Universal Definition of Myocardial Infarction
Rise and/or Fall of troponin with at least one value >99th %ile
ALL elevations >99th % are Myocardial Injury: must have rise/fall to be ACUTE
NOT ALL elevations are MI
August 14, 2023
ADD-140731-USA-EN |

## Slide 9: Numerous reasons for non-MI troponin elevations
Hoeller et al.; BMJ 2013
August 14, 2023

## Slide 10: Non-MI troponin T elevations: Skeletal muscle disorders (SMD) can cause elevations in cTnT not due to cardiac disease1
hs-cTnT elevations are common in patients with active chronic non-inflammatory myopathy and myositis, but not other SMDs
hs-cTnT-Elecsys concentrations were significantly higher in patients with SMD versus controls
hs-cTnI concentrations were mostly similar
In patients with active chronic non-inflammatory myopathy and myositis, cTnI is the preferred analyte for assessing cardiac health in general and presence of AMI
Mueller et al.; Cardiac Troponins and Skeletal Muscle Disease; 10.1161/CIRCULATIONAHA.121.058489
August 14, 2023
ADD-140731-USA-EN |

## Slide 11: High sensitivity vs. contemporary troponin assays
Alinity i STAT high sensitivity Troponin-i
August 14, 2023
ADD-140731-USA-EN |
11

## Slide 12: How are hs troponin assays defined and reported?1
IN 2012, A TASK FORCE CREATED BY THE INTERNATIONAL FEDERATION OF CLINICAL CHEMISTRY (IFCC) PROPOSED THAT HSTN ASSAYS BE DEFINED USING TWO PARAMETERS:
The total imprecision (coefficient of variation) at the 99th %ile should be ≤10%
Measurable concentrations below the 99th %ile should be attainable with an assay at a concentration value above the assay’s limit of detection (LOD) for at least 50% (and ideally >95%) of healthy individuals
In 2018, the Task Force expanded on this point by requiring both men and women individually attain measurable concentrations: at least 50% measurable concentrations above the assay’s LoD
IFCC also recommend that hs-cTn is reported in whole numbers, using ng/L without decimal points
It is crucial that clinicians are educated on the change in units
Wu et al, Clin Chem 2018
August 14, 2023
ADD-140731-USA-EN |
Notes: Questions 4 & 5: identify differences between ng/mL vs ng/L and explain why hsTnI results are reported in whole numbers

## Slide 13: Limitations of contemporary troponin assays
CONTEMPORARY (NON HS) TROPONIN I AND T ASSAYS ARE STILL IN WIDESPREAD USE IN US
Very reliable for detecting troponin levels > 99th percentile of an apparently healthy reference population1 but may not be able to detect elevations as quickly or precisely as high sensitivity assays2
POTENTIALLY REQUIRE EXTENDED TIME TO DETECT TROPONIN WHICH MAY LEAD TO:
Delay in the diagnosis of MI
Delay in treatment for patients ultimately diagnosed with MI: “time is muscle”
Particularly vulnerable are those with lower circulating levels of troponin
e.g. Women and Early presenters (present with <2-3 hours of symptoms)3
HIGH SENSITIVITY ASSAYS CAN DETECT ELEVATED CTNI (> 99TH %) WITHIN 3 HRS AFTER THE ONSET OF CHEST PAIN4
Jarolim P. Clin Chem Lab Med. 2015; 53(5): 635-652
Wu et al, Clin Chem 2018
Thygesen et al, JACC 2018 Fourth Universal Definition of Myocardial Infarction
Thygesen K, et al. EHJ 2012;33(18):2252-2257
August 14, 2023
ADD-140731-USA-EN |

## Slide 14: Limitations of contemporary troponin assays
CONTEMPORARY (NON HS) TROPONIN I AND T ASSAYS
Cannot reliably report levels well below the 99th %ile
Result may appear similar whether true levels of troponin were very low or rising rapidly
For example, where presentation = 0 hrs
“<28 ng/L” and 3 hrs later “<28 ng/L” could represent true values such as
0 hrs 5ng/L, 3 hrs 6 ng/L or
0 hrs 5ng/L, 3 hrs 16 ng/L
Clinicians are aware of limitations of current paradigms and may keep patients for observation and testing for extended periods of time, leading to increased length of stay and cost1
ABBOTT Alinity HIGH SENSITIVITY TROPONIN I ASSAY CAN REPORT LEVELS OF CTNI AS LOW AS 2.7 NG/L (LOQ)
Thygesen et al, JACC 2018 Fourth Universal Definition of Myocardial Infarction
Alinity i STAT High Sensitivity Troponin-I Package Insert: H14937R02
August 14, 2023
ADD-140731-USA-EN |

## Slide 15: Contemporary troponin1,2,3,4
1. Apple, F.S., et al. Clin Chem, 2012; 58(1) :54–61
2. deFilippi, C., et al. JAMA, 2010; 304(22): 2494-2502
3. Saunders, J. et al. Circulation, 2011; 123: 1367-1376
4. Atherton, JJ., et al. JACC Cardiovasc Imaging, 2010; 3: 421-428
Troponin Level
Frequency
99th Percentile
Limit of Detection
Not Measureable!
10% CV
Representative cartoon for contemporary assays
Notes: So Why is it so important to have measurable values above the LoD?
The higher the sensitivity of the assay, the more frequently Tn measurements in healthy subjects will be detected. Detected is defined by any measurement above the lower limit of detection (LoD), i.e., noise of the assay.1-2
Studies have suggested that elevated troponin values, even below the upper limit of normal (i.e., 99thpercentile) of a healthy reference population are of prognostic importance and are linked to worse clinical outcomes.3-6
References:
1. Apple FS., Clin Chem. 2009; 57(7): 1303-13006
2. Apple, F.S., et al. Clin Chem, 2012; 58(1) :54–61
3 .deFilippi, C., et al. JAMA, 2010; 304(22): 2494-2502
4. Saunders, J. et al. Circulation, 2011; 123: 1367-1376
5. Atherton, JJ., et al. JACC Cardiovasc Imaging, 2010; 3: 421-428
In this representative cartoon, contemporary sensitive Troponin assays have higher LoD and in some cases are close to the 99thile resulting in not picking up measurable values of Troponin in a normal population.

## Slide 16: High sensitivity troponin
Troponin Level
Frequency
99th Percentile
Cotemporary LoD
10% CV of Contemporary Tnl
Not Measureable!
High-Sensitivity LoD
10% CV of hsTnl
Representative cartoon for high-sensitivity assays

## Slide 17: Contemporary vs hs troponin assays
Previous generation and contemporary troponin assays cannot detect troponin increases as quickly as high-sensitivity troponin assays
August 14, 2023
ADD-140731-USA-EN |

## Slide 18: Alinity hsTnI meets IFCC criteria for hs assay
<5% CV
near medical decision points
Alnity i STAT_hsTnI Precision at 99th Percentile |  | 
 | Mean (ng/L) | Coefficient of Variation (CV)**
Sample 5 | 13.9 | 2.7%
Sample 6 | 15.1 | 4.3%
Sample 7 | 18.2 | 4.5%
Sample 8 | 20.2 | 3.7%
Sample 9 | 36.9 | 3.2%
≥50%
Percent detected using ARCHITECT hsTnI in healthy population studies1
**Includes within-run and between-run variability.
Chart: 
| Category 1
Overall | 0.753
Males | 0.873
Females | 0.634
Alinity i STAT High Sensitivity Troponin-I Package Insert: H14937R02
Abbott data on file

## Slide 19: ARCHITECT TnI and hsTnI_STAT and Alinity i hsTnI_STAT
Improved sensitivity with lower LoD
Improvement in the 10% CV concentration without significant change to 99th percentile
***Alinity i High Sensitivity Troponin-I Package Insert: H14937R02
*ARCHITECT STAT TnI Package Insert: G10467R11
**ARCHITECT STAT High Sensitivity Troponin-I Ho3755/R01
Data on file at Abbott
Apple FS, Collinson PO. Analytical characteristics of high-sensitivity cardiac troponin assays. Clinical Chemistry. 2012;58(1):54-61
 | ARCHITECT STAT TnI (LN 2K41) | ARCHITECT hsTnI_STAT (LN 2R98) | Alinity i hsTnI_STAT (LN 4Z21)
 | pg/ml (ng/ml) | Ng/L | Ng/L
Analytical Sensitivity/LoD | 10* | 1.7** | 0.9***
LoQ | N/A | 3.5 | 2.7
10% CV1 | 32 | 4.6 | > 3.5
99th percentile (overall) | 28 | 28 | 27
99th percentile (male) | 33 | 35 | 35
99th percentile (females) | 13 | 17 | 14
% Detectable above LoD2 | <50% of normal | >50% of normal | >50% of normals
LoB | N/A | 0.9 | 0.0
Sample tube type | Li Heparin Plasma | K2 EDTA Plasma | Li Heparin Plasma
August 14, 2023
ADD-140731-USA-EN |
Need for sex specific cutoffs
Notes: The LoB represents the 95th percentile from n ≥ 60 replicates of zero-analyte samples.
b The LoD represents the lowest concentration at which the analyte can be detected with 95% probability based on n ≥ 60 replicates of low-analyte level samples.
c The LoQ presented in the table is in alignment with the low end of the AMI for the ARCHITECT STAT High Sensitivity Troponin-I assay. The observed LoQ on the ARCHITECT i System was 2.3 ng/L (2.3 pg/mL). This LoQ is defined as the lowest concentration at which a maximum allowable precision of 20 %CV was met and was determined from n ≥ 60 replicates of low-analyte level samples.
EDTA was used to ARCH b/c that was the only data we had for sample type studies
Li Hep is the most used tube in ER
On going efforts to add tube types, but decision was made with Li Hep as to not jeopardize approval

## Slide 20: Interfering substances
Alinity i STAT high sensitivity Troponin-I
August 14, 2023
ADD-140731-USA-EN |
20

## Slide 21: Interfering substances
Alinity i High Sensitivity Troponin-I Package Insert: H14937R02
August 14, 2023

## Slide 22: Hemolysis interference is an important consideration for choice of hsTnI assay
Hemolysis is known to cause interference in both TnT and TnI assays1
Effects become even more important with the emergence of more sensitive troponin assays and the increased reliance on interpretation of small absolute changes1
Hemoglobinemia and hyperbilirubinemia may produce either false-positive or false-negative troponin levels, depending on the assay used and the form or subunit of troponin measured2
Florkowski C, Wallace J, Walmsley T, George P. The effect of hemolysis on current troponin assays--a confounding preanalytical variable? Clinical Chemistry. 2010;56(7):1195-1197.
Means Jr. RT( 1 ), Masimasi N. Elevated troponin levels associated with hemolysis. American Journal of the Medical Sciences. 330(4):201-203.
August 14, 2023
ADD-140731-USA-EN |

## Slide 23: Hemolysis interference with troponin assays
Hemolysis is seen more frequently in samples from ED (12.4%) compared to routine samples (1.6%)1
Susceptibility to hemolysis2:
  can provide inaccurate results
  require a redraw
  Increase time to diagnose
  create lack of confidence in results
Burns ER, Yoshikawa; Laboratory Medicine; 2022; 33(5):378-380
Harley K etal. Clin Chem Lab Med; 2021

## Slide 24: Emergency department samples have a higher incidence of hemolysis compared to routine
Hemolysis in samples drawn in the Emergency Department (ED) is significantly more frequent than samples drawn in routine situations1
Evaluation of hemolysis in ED samples requesting high sensitivity troponin measurement2
  7% samples had Hb >0.1g/dL (competitor claim)
  Only 0.5% samples had Hb>0.5g/dL
Burns ER, Yoshikawa N. Hemolysis in serum samples drawn by emergency department personnel versus laboratory phlebotomists. Laboratory Medicine. 33(5):378-380
Brunel V, Larson T, Peschanski N, Cauliez B. Evaluation of hemolysis in emergency department samples requesting high sensitivity troponin T measurement. Annals Of Clinical Biochemistry. 2012;49(Pt 5):509-510
August 14, 2023
ADD-140731-USA-EN |
Notes: The reference to 7% is clearly stated in the letter… “According to the hsTnT manufacturer’s recommendations, haemolysis would impact up to 7% of our emergency department samples.” This is because as Roche state in their pack insert “Falsely depressed results are obtained when using samples with hemoglobin concentrations > 0.1 g/dL” this gives a massive change of 20%.... ”Criterion: Recovery within ± 20 % of initial value at troponin T concentrations”. Note that the ARCHITECT hsTroponin assay says haemolysis interference occurs above 0.5 g/dL (500mg/dL) with a smaller 10% change criteria.
In Brunel the prevalence of haemolysis above 0.1 g/dL was found to be 7% (in the graphic 4.6% + 2% + 0.5% = 7.1%) , and above 0.5g/dL was found to be 0.5%, this is clear from the graphic.
So given the representative haemolysis prevalence data in the Brunel study …. 7% or 1in 14 samples from the ED are unsuitable because of hemolysis for Roche hsTnT vs only 0.5% or 1in 200 samples from ED for the Abbott hsTnI assay. If an ED sent 70 samples a day for troponin, then 5 times a day (70/14) the lab would ring ED and say take another sample for Roche vs only ring the ED once every 3 days for Abbott (70/200 = 1/3).

## Slide 25: Performance and specifications
Alinity i STAT high sensitivity Troponin-I
August 14, 2023
ADD-140731-USA-EN |
25

## Slide 26: Alinity i hsTnI methodology
This assay is an automated, two-step immunoassay for the quantitative determination of cTnI in human plasma (lithium heparin) using chemiluminescent microparticle immunoassay (CMIA) technology.
Sample and anti-troponin I antibody-coated paramagnetic microparticles are combined and incubated.
The cTnI present in the sample binds to the anti-troponin I coated microparticles.
The mixture is washed.
Anti-troponin I acridinium-labeled conjugate is added to create a reaction mixture and incubated.
Following a wash cycle, Pre-Trigger and Trigger Solutions are added.
The resulting chemiluminescent reaction is measured as a relative light unit (RLU). There is a direct relationship between the amount of cTnI in the sample and the RLU detected by the system optics.
Alinity i High Sensitivity Troponin-I Package Insert: H14937R02
August 14, 2023

## Slide 27: Alinity i hsTnI methodology
August 14, 2023

## Slide 28: Alinity i STAT High Sensitivity Troponin-I assay
Alinity i High Sensitivity Troponin-I Package Insert: H14937R02
Alinity Systems Operation Manual 96211-119
Alinity i hsTnI Control Package Insert
Method & Format | 2-Step CMIA Assay with STAT protocol capability
Result Unit | ng/L; alternative SI unit = pg/mL
Reportable Interval | AMI = 2.7 – 3600.0 ng/L EMI = 3600.0 – 60,000.0 ng/L
Time to First Result1 | STAT 18 min
Control Levels2 | Low: 20 ng/L; 12.0 – 28.0 ng/L Medium: 200 ng/L; 120.0 – 280.0 ng/L
LoQ | 2.7 ng/L
LoD | 0.9 ng/L
LoB | 0.0 ng/L
99th Percentile | Overall = 27 ng/L Females = 14 ng/L Males = 35 ng/L
Specimen Type | Plasma: Lithium Heparin
Sample Volume | 160 uL
August 14, 2023
ADD-140731-USA-EN |
Notes: LoQ: is used as lowest cutoff.
LoD: used to qualify assay to be high sensitivity, >50% healthy population

## Slide 29: Specimen specifications
Only plasma tubes containing dried lithium heparin have been validated for use with this assay.
Other specimen types and collection tube types have not been verified with this assay.
Performance has not been established for the use of cadaveric specimens or the use of body fluids other than human plasma.
The instrument does not provide the capability to verify specimen type. It is the responsibility of the operator to verify that the correct specimen types are used in the assay.
Alinity i High Sensitivity Troponin-I Package Insert: H14937R02
Specimen Type | Temperature | Max Storage Time | Special Instructions
Plasma – Lithium heparin with or without separator | Room temperature (15-30oC) | 8 hours | Specimens may be stored on the red blood cells.
August 14, 2023
ADD-140731-USA-EN |